URGENT Field Safety Notice

Philips Respironics – Sleep and Respiratory Care

Notification To DMEs/Distributors/Prescribers

Amara View Minimal Contact Full-Face Mask

DreamWear Full Face Mask

DreamWisp Nasal Mask with Over the Nose Cushion

Wisp Nasal Masks and Wisp Youth Nasal Masks

Therapy Mask 3100 NC/SP

This document contains important information for the continued safe and proper use of your equipment

- Please review the following information with all members of your staff who need to be aware of the contents of this communication.  It is important to understand the implications of this communication.

- Please retain a copy with the equipment Instruction for Use.

This document contains important information for the continued safe and proper use of your equipment

- Please review the following information with all members of your staff who need to be aware of the contents of this communication.  It is important to understand the implications of this communication.

- Please retain a copy with the equipment Instruction for Use.

August 29, 2022

Dear Customer,

The following Philips Respironics patient interface devices (face and nasal masks) - Amara View Minimal Contact Full-Face Mask, DreamWear Full Face Mask, DreamWisp Nasal Mask with Over the Nose Cushion, Wisp Nasal Masks, Wisp Youth Nasal Masks, and Therapy Mask 3100 NC/SP - contain magnets.

This Field Safety Notice is intended to inform you that Philips Respironics is updating its existing ‘Contraindications’ and ‘Warnings’ to the following:

Contraindication:  Use of the mask is contraindicated for patients and their household members, caregivers, and bed partners that may be in close vicinity to patients using the masks, that have implanted devices that may be affected by magnets, including but not limited to:

- Pacemakers

- Implantable cardioverter defibrillators (ICD)

- Neurostimulators

- Magnetic metallic implants/electrodes/valves placed in upper limbs, torso, or higher (i.e. neck and head)

- CSF (cerebral spinal fluid) shunts (e.g., VP (ventriculo peritoneal) shunt)

- Aneurysm clips

- Embolic coils

- Intracranial aneurysm intravascular flow disruption devices

- Metallic cranial plates, screws, burr hole covers, and bone substitute devices

- Metallic splinters in the eye

- Ocular implants (e.g., glaucoma implants, retinal implants)

- Certain contact lenses with metal

- Implants to restore hearing or balance that have an implanted magnet (such as cochlear implants, implanted bone conduction hearing devices, and auditory brainstem implants)

- Magnetic denture attachments

- Metallic gastrointestinal clips

- Metallic stents (e.g., aneurysm, coronary, tracheobronchial, biliary)

- Implantable ports and pumps (e.g., insulin pumps)

- Hypoglossal Nerve Stimulators

- Devices labeled as MR (magnetic resonance) unsafe

- Magnetic metallic implants not labeled for MR or not evaluated for safety in a magnetic field

Warning:  Magnets with a magnetic field strength of 400 mT are used in the mask. With the exception of the devices identified in the contraindication, ensure the mask is kept at least 6 inches (approx. 15.24 cm) away from any other medical implants or medical devices that can be impacted by the magnetic fields to avoid possible effects from localized magnetic fields.  This includes household members, caregivers, and bed partners that may be in close vicinity to patients that use the masks.

- What the problem is and under what circumstances it can occur

- The affected masks contain magnets which can potentially affect the functioning and/or induce the movement/dislocation of medical implants or medical devices that can be impacted by the magnetic fields.

- See the Contraindication and Warning in this notice for additional details.

- Describe the hazard/harm associated with the issue

- With the exception of the devices in the contraindication, if the mask magnets are placed less than 6 inches (approx. 15.24 cm) away from a metallic implant or device the magnets may cause the device to not perform as intended, which may result in a serious injury.

- As of August 29, 2022, there have been fourteen (14) reports of patients suggesting that the mask magnets have impacted their medical devices which include: pacemaker interference, pacemaker failure leading to replacement, need of shunt adjustment, resetting of automatic implantable cardioverter defibrillator (AICD), seizures, defibrillator shutting off periodically, arrhythmia, irregular blood pressure, change in heartbeats, and cognitive issues.

- Affected products and how to identify them

Images of the affected masks are shown below.  The magnetic headgear clips on these masks are circled.  These images may be used as reference to determine if your or your patient's mask also uses magnetic headgear clips.  These masks are intended to provide an interface for application of CPAP or bi-level therapy to patients.  The part numbers for the affected masks and mask components that contain magnets are listed in Figure 6.

Figure 1:  Amara View Full Face Mask

Figure 1:  Amara View Full Face Mask

Figure 2:  DreamWisp Nasal Mask

Figure 2:  DreamWisp Nasal Mask

Figure 3: DreamWear Full Face Mask

Figure 3: DreamWear Full Face Mask

Figure 4:  Wisp and Wisp Youth Nasal Mask

Figure 4:  Wisp and Wisp Youth Nasal Mask

Figure 5:  Therapy Mask 3100 NC/SP

Figure 5:  Therapy Mask 3100 NC/SP

- Actions that should be taken by the customer in order to prevent risks for patients

As a DME, distributor, or prescriber of these masks, you must take the following action:

  - Provide this Field Safety Notice and copies of the included patient version of the Field Safety Notice and FAQs document to all those who must be aware within your organization, and to any organizations where the affected devices have been transferred.

  - Provide the given copies of the included patient version of the Field Safety Notice and FAQs document to all patients using the affected devices*.

*NOTE: Philips Respironics sells these products through DMEs/distributors and prescribers, and we may not have the contact information to reach all patients.

  - Fill out the Business Reply Form included in this letter and return it by email per the instructions in the form.  This form serves as official acknowledgement that you have received this correction letter and completed the actions required for this medical device correction.

  - Patients who have an implanted device where use of the magnetic mask is contraindicated should follow the following steps:

- If they have an alternative non-magnetic mask available at home they should stop using the magnetic mask immediately and switch to the non-magnetic mask for continued therapy and inform their healthcare professional of this switch.

- If they do not have an alternative non-magnetic mask available at home, they should get in touch with their healthcare professional immediately to arrange supply of an alternative mask that can be used to continue their therapy. Patients should consult with their healthcare professional on whether therapy should be stopped while alternative masks are provided.

- Patients should properly dispose of their magnetic mask once an alternative mask has been obtained.

- If patients are unsure whether their implanted device is affected, they should speak to their healthcare professional.

  - If patients, and household members, caregivers, and bed partners that may be in close vicinity to patients, do not have implanted medical devices, or metallic splinters in their eyes, then no action related to patients is needed.

  - Contact Philips Respironics customer service to learn more about non-magnetic mask options.

  - Household members, caregivers, and bed partners with a medical implant/device must ensure the mask is kept at least 6 inches (approx. 15.24 cm) away from the medical implant(s)/device(s).

- Describe the actions planned by Philips Respironics to correct the problem

Philips Respironics is updating its existing contraindication and warning as provided in this field safety notice.

- If you need any further information or support concerning this issue, please contact your local Philips Respironics representative: UK Philips Customer Care Service Centre on 0870 532 9741.

This notice has been reported to the appropriate Regulatory Agencies.  Adverse reactions or quality problems experienced with the use of this product may be reported to Philips Respironics or the local competent authority Medicines and Healthcare products Regulatory Agency (MHRA). Philips Respironics regrets any inconvenience caused by this problem.

Sincerely,

Thomas Fallon

Head of Quality

Philips Respironics

__________________________________________________________________________________________________

Devices Impacted by This Issue - Figure 6:

| Amara View Minimal Contact Full-Face Mask |  |  | Wisp and Wisp Youth Nasal Mask |  |
| --- | --- | --- | --- | --- |
| Part # | Description |  | Part # | Description |
| 1090602 | S Amara View Mask w/HGR |  | 1094050 | WISP MASK, CLEAR FRAME, W/ HDGR, DOM |
| 1090603 | M Amara View Mask w/HGR |  | 1094051 | WISP MASK, FABRIC FRAME, W/MAG CLIPS DOM |
| 1090604 | L Amara View Mask w/HGR |  | 1094053 | WISP MASK, USA, FABRIC FRAME, W/ HDGR |
| 1090622 | S Amara View Mask w/ HGR, DOM |  | 1094060 | WISP MASK, CLEAR FRAME, W/ HDGR, INTL |
| 1090623 | M Amara View Mask w/ HGR, DOM |  | 1094061 | WISP MASK, FABRIC FRAME, W/ HDGR, INTL |
| 1090624 | L Amara View Mask w/ HGR, DOM |  | 1094062 | WISP MASK, CLEAR FRAME, W/ HDGR, INTL SE |
| 1090641 | XS Amara View Mask w/HGR, JPN |  | 1094070 | WISP MASK, CLEAR FRAME, W/ HDGR, JPN SE |
| 1090642 | S Amara View Mask w/HGR, JPN |  | 1094071 | P, WISP MASK, FABRIC FRAME, W/ HDGR, JPN |
| 1090643 | M Amara View Mask w/HGR, JPN |  | 1094072 | S/M, WISP MASK FABRIC FRAME, W/ HDGR JPN |
| 1090644 | L Amara View Mask w/HGR, JPN |  | 1094073 | L, WISP MASK, FABRIC FRAME, W/ HDGR, JPN |
| 1090651 | XS Amara View Mask SE w/HGR, JPN |  | 1109298 | WISP YOUTH MASK, FABRIC, W/MAG CLIPS DOM |
| 1090652 | S Amara View Mask SE w/HGR, JPN |  | 1113954 | XL, Wisp Mask, Fabric Frame, w/ HDGR, JP |
| 1090653 | M Amara View Mask SE w/HGR, JPN |  | 1117119 | MW, Wisp Mask, Fabric Frame, w/ HDGR, JP |
| 1090654 | L Amara View Mask SE w/HGR, JPN |  | 1118060 | P, WISP MASK, CLEAR FRAME, W/ HDGR, DOM |
| 1090662 | S Amara View Mask w/HGR, INTL |  | 1118061 | S/M, WISP MASK CLEAR FRAME, W/ HDGR DOM |
| 1090663 | M Amara View Mask w/HGR, INTL |  | 1118062 | L, WISP MASK, CLEAR FRAME, W/ HDGR, DOM |
| 1090664 | L Amara View Mask w/HGR, INTL |  | 1118063 | XL, WISP MASK, CLEAR FRAME, W/ HGR, DOM |
| 1090670 | Amara View Mask w/HGR, FitPack |  | 1118064 | P, WISP MASK, FABRIC FRAME, W/ HDGR,DOM |
| 1090696 | RP – Amara View Headgear, L |  | 1118065 | S/M, WISP MASK FABRIC FRAME W/ HDGR,DOM |
| 1090697 | RP-Amara View Headgear, Standard |  | 1118066 | L, WISP MASK, FABRIC FRAME, W/ HDGR,DOM |
| 1135080 | S AMARA VIEW MASK SE W/ HGR |  | 1118067 | XL, WISP MASK, FABRIC FRAME W/ HGR, DOM |
| 1135081 | M AMARA VIEW MASK SE W/ HGR |  | 1120434 | Wisp/DreamWisp Magnetic Headgear Clip,RP |
| 1135082 | L AMARA VIEW MASK SE W/ HGR |  | 1137360 | Wisp Nasal Lab Mask, Fabric Frame, DOM |
| 1136626 | Amara View Magnet Clip |  | 1137361 | Wisp Nasal Lab Mask, Clear Frame, DOM |
| 1090670RTL | AMARA VIEW MASK W/HGR, FITPACK-RTL |  | 1137366 | S/M, Wisp Lab Mask, Fabric Frame, DOM |
| 1136626RTL | Amara View Full Face Magnetic Clips |  | 1137367 | L, Wisp Nasal Mask, Fabric Frame, DOM |
|  |  |  | 1137368 | S/M, Wisp Nasal Mask, Clear Frame, DOM |
|  |  |  | 1137369 | L, Wisp Nasal Mask, Clear Frame, DOM |
| DreamWear Full Face Mask |  |  | 1137370 | COMBOPACK, WISP LAB MASK ,DOM |
| Part # | Description |  | 1138989 | P, WISP MASK,CLEAR FRAME, W/HDGR,1118060 |
| 1133344 | S, DreamWear UTN Nsl/Full, Combo Pack |  | 1138991 | L, WISP MASK, CLEAR W/ HDGR, DOM 1118062 |
| 1133345 | M, DreamWear UTN Nsl/Full, Combo Pack |  | 1138992 | XL,WISP MASK,CLEAR FRAME,W/ HGR,1118063 |
| 1133370 | S, DreamWear Full, Sm & Med Frm, INTL |  | 1138993 | P, WISP MASK,FABRIC FRAME,W/HDGR,1118064 |
| 1133371 | M, DreamWear Full, Sm & Med Frm, INTL |  | 1138994 | S/M,WISP MASK FABRICFRAME W/HDGR,1118065 |
| 1133372 | L, DreamWear Full, Sm & Med Frm, INTL |  | 1138995 | L, WISP MASK,FABRIC FRAME,W/HDGR,1118066 |
| 1133373 | MW, DreamWear Full, Sm & Med Frm |  | 1138996 | XL,WISP MASK,FABRIC FRAME W/HGR, 1118067 |
| 1133375 | S, DreamWear Full, Sm & Med Frm, GBL |  | 1139009 | S/M,WISPMASK CLEAR FRAME,W/ HDGR 1118061 |
| 1133376 | M, DreamWear Full, Sm & Med Frm, GBL |  | 1094051RTL | WISP MASK, FABRICFRM,W/MAG CLIPS DOM-RTL |
| 1133377 | L, DreamWear Full, Sm & Med Frm, GBL |  | 1148366 | DreamWisp/Wisp Magnetic Clips - RTL |
| 1133378 | MW, DreamWear Full, S. & M. Frm, GBL |  |  |  |
| 1133380 | S, DreamWear Full, M. Frm W/ HGR, GBL |  |  |  |
| 1133381 | M, DreamWear Full, Med Frm W/ HGR |  | DreamWisp Nasal Mask with Over the Nose Cushion |  |
| 1133382 | L, DreamWear Full, Med Frm W/ HGR |  | Part # | Description |
| 1133383 | MW, DreamWear Full, Med Frm W/ HGR |  | 1137916 | Fitpk DreamWisp Nsl, MED Con W/HGR GBL |
| 1133385 | S, DreamWear Full, Sm Frm W/ HGR, GBL |  | 1137918 | FitPk DreamWisp Nsl, MED Con W/O HGR GBL |
| 1133386 | M, DreamWear Full, Sm Frm W/ HGR |  | 1137921 | FitPk, DreamWisp Nsl, MED Con W/ HGR INT |
| 1133387 | L, DreamWear Full, Sm Frm W/ HGR, GBL |  | 1137922 | P, DreamWisp Nsl, Med Con W/O HGR GBL |
| 1133388 | MW, DreamWear Full, Sm Frm W/ HGR, |  | 1137923 | S, DreamWisp Nsl, Med Con W/O HGR GBL |
| 1133390 | S, DreamWear Full, Lg Frm W/ HGR, GBL |  | 1137924 | M, DreamWisp Nsl, Med Conn W/O HGR GBL |
| 1133391 | M, DreamWear Full, Lg Frm W/ HGR, GBL |  | 1137925 | L, DreamWisp Nsl, Med Con W/O HGR GBL |
| 1133392 | L, DreamWear Full, Lg Frm W/ HGR, GBL |  | 1137926 | XL, DreamWisp Nsl, Med Con W/O HGR GBL |
| 1133393 | MW, DreamWear Full, Lg Frm W/ HGR |  | 1137927 | P, DreamWisp NSL, SM CON, W/HGR, GBL |
| 1133395 | DreamWear Full, Demo Pack |  | 1137928 | S, DreamWisp NSL, SM CON, W/HGR, GBL |
| 1133400 | FitPack, DreamWear Full, Med Frame |  | 1137929 | M, DreamWisp NSL, SM CON, W/HGR, GBL |
| 1133401 | DreamWear Full, Med Frm W/O HGR, GBL |  | 1137930 | L, DreamWisp NSL, SM CON, W/HGR, GBL |
| 1133405 | S, DreamWear Full, Med Frm W/O HGR |  | 1137931 | XL, DreamWisp NSL, SM CON, W/HGR, GBL |
| 1133406 | M, DreamWear Full, Med Frm W/O HGR |  | 1137932 | P, DreamWisp Nsl, Med Con, W/HGR, GBL |
| 1133407 | L, DreamWear Full, Med Frm W/O HGR |  | 1137933 | S, DreamWisp Nsl, Med Con, W/HGR, GBL |
| 1133408 | MW, DreamWear Full, Med Frm W/O |  | 1137934 | M, DreamWisp Nsl, Med Con, W/HGR, GBL |
| 1133410 | S, DreamWear Full, Sm Frm W/O HGR |  | 1137935 | L, DreamWisp Nsl, Med Con, W/HGR, GBL |
| 1133411 | M, DreamWear Full, Sm Frm W/O HGR |  | 1137936 | XL, DreamWisp Nsl, Med Con, W/HGR, GBL |
| 1133412 | L, DreamWear Full, Sm Frm W/O HGR |  | 1137937 | P, DreamWisp NSL, SM CON, W/HGR, INT |
| 1133413 | MW, DreamWear Full, Sm Frm W/O HGR |  | 1137938 | S, DreamWisp NSL, SM CON, W/HGR, INT |
| 1133415 | S, DreamWear Full, Lg Frm W/O HGR, GBL |  | 1137939 | M, DreamWisp NSL, SM CON, W/HGR, INT |
| 1133416 | M, DreamWear Full, Lg Frm W/O HGR |  | 1137940 | L, DreamWisp NSL, SM CON, W/HGR, INT |
| 1133417 | L, DreamWear Full, Lg Frm W/O HGR, GBL |  | 1137941 | XL, DreamWisp NSL, SM CON, W/HGR, INT |
| 1133418 | MW, DreamWear Full, Lg Frm W/O HGR |  | 1137942 | P, DreamWisp Nsl, Med Con, W/HGR INT |
| 1133423 | S/M, DreamWear Full, Sm Frm w/ HGR |  | 1137943 | S, DreamWisp Nsl, Med Con, W/HGR INT |
| 1133424 | S/M, DreamWear Full, Med Frm w/ HGR |  | 1137944 | M, DreamWisp Nsl, Med Con, W/HGR INT |
| 1133425 | S/M, DreamWear Full, Lg Frm w/ HGR |  | 1137945 | L, DreamWisp Nsl, Med Con, W/HGR INT |
| 1133427 | W, DreamWear Full, Med Frm W/ HGR |  | 1137946 | XL, DreamWisp Nsl, Med Con, W/HGR INT |
| 1133430 | S, DreamWear Full Cushion, RP |  | 1137947 | P, DreamWisp Nsl, S-M-L Con, W/HGR, JPN |
| 1133431 | M, DreamWear Full Cushion, RP |  | 1137948 | S, DreamWisp Nsl, S-M-L Con, W/HGR, JPN |
| 1133432 | L, DreamWear Full Cushion, RP |  | 1137949 | M, DreamWisp Nsl, S-M-L Con, W/HGR, JPN |
| 1133433 | MW, DreamWear Full Cushion, RP |  | 1137950 | L, DreamWisp Nsl, S-M-L Con, W/HGR, JPN |
| 1133434 | SW, DreamWear Full Cushion, RP |  | 1137951 | XL, DreamWisp Nsl, S-M-L Con, W/HGR, JPN |
| 1133435 | S, DreamWear Full Cushion, RP, JPN |  | 1137953 | LABPK DreamWisp ALL CON S/M/L CUSH W/HGR |
| 1133436 | M, DreamWear Full Cushion, RP, JPN |  | 1137954 | P/XL LABPK DreamWisp LAB PACK, S/L con |
| 1133438 | W, DreamWear Full Cushion, RP, JPN |  | 1137955 | Combo LabPK DreamWisp |
| 1133450 | DreamWear Full Headgear, RP |  | 1137964 | DreamWisp HEADGEAR, RP |
| 1133454 | Headgear Magnet Clip, RP |  | 1137965 | DreamWisp LARGE HEADGEAR, RP |
| 1135275 | S, DreamWear Full, Sm and Med Frm |  | 1137966 | DreamWisp RS HEADGEAR, RP |
| 1135276 | M, DreamWear Full, Sm and Med Frm |  | 1139332 | P, DreamWisp Nsl, S-M Con, W/HGR KR |
| 1135277 | SW, DreamWear Full,Sm and Med Frm |  | 1139333 | S, DreamWisp Nsl, S-M Con, W/HGR KR |
| 1135278 | MW, DreamWear Full,Sm and Med Frm |  | 1139334 | M, DreamWisp Nsl, S-M Con, W/HGR KR |
| 1135327 | S, DreamWear Full, Sm & Med Frm, KR |  | 1139335 | L, DreamWisp Nsl, S-M Con, W/HGR KR |
| 1135388 | M, DreamWear Full, Sm & Med Frm, KR |  | 1139336 | XL, DreamWisp Nsl, S-M Con, W/HGR KR |
| 1135389 | SW, DreamWear Full, Sm & Med Frm, KR |  | 1144344 | DreamWisp RS HEADGEAR-Lab, RP |
| 1135390 | MW, DreamWear Full, Sm & Med Frm |  | 1144345 | DreamWisp LARGE HEADGEAR-Lab, RP |
| 1137430 | Fitpack, DreamWear Full Lab,Med Frm |  | 1144346 | DreamWisp HEADGEAR - Lab, RP |
| 1137431 | Ext Sizepack,DreamWear Full Lab Msk |  | 1144502 | FitPack, DreamWisp, USA |
| 1137437 | S, DreamWear Full Lab Msk,Med Frame |  | 1148409 | DreamWisp Headgear (RS)- RTL |
| 1137438 | S, DreamWear Full Lab Mask,Sm Frame |  | 1148410 | DreamWisp HEADGEAR, RP- RTL |
| 1137439 | M, DreamWear Full Lab Msk,Med Frame |  | 1148411 | DreamWisp LARGE HEADGEAR, RP- RTL |
| 1137440 | COMBOPACK, DREAMWEAR FULL LAB |  | 1148913 | P, DreamWisp Nsl, S-M Con, W/HGR CN |
| 1140235 | S, DreamWear Full, Med Frm W/ HGR |  | 1148914 | S, DreamWisp Nsl, S-M Con, W/HGR CN |
| 1143130 | FitPack, DreamWear Full, Med Frame |  | 1148915 | M, DreamWisp Nsl, S-M Con, W/HGR CN |
| 1144347 | DreamWear Full Headgear - Lab, RP |  | 1148916 | L, DreamWisp Nsl, S-M Con, W/HGR CN |
| 1146593 | S, DreamWear Full Cushion, RP - RTL |  | 1148917 | XL, DreamWisp Nsl, S-M Con, W/HGR CN |
| 1146594 | M, DreamWear Full Cushion, RP - RTL |  | 1137942AP | P, DREAMWISP NSL, MED CON, W/HGR, AP |
| 1146595 | L, DreamWear Full Cushion, RP - RTL |  | 1137942CE | P, DREAMWISP NSL, MED CON, W/HGR, CE |
| 1146596 | MW, DreamWear Full Cushion, RP - RTL |  | 1137942RC | P, DREAMWISP NSL, MED CON, W/HGR, RC |
| 1146597 | DreamWear Full Headgear, RP - RTL |  | 1137943AP | S, DREAMWISP NSL, MED CON, W/HGR, AP |
| 1148363 | DreamWear Full Face Magnet Clip - RTL |  | 1137943CE | S, DREAMWISP NSL, MED CON, W/HGR, CE |
| 1133375RTL | DreamWear Full Face Mask Kit, Small |  | 1137943RC | S, DREAMWISP NSL, MED CON, W/HGR, RC |
| 1133376RTL | DreamWear Full Face Mask Kit, Medium |  | 1137944AP | M, DREAMWISP NSL, MED CON, W/HGR, AP |
| 1133377RTL | DreamWear Full Face Mask Kit, Large |  | 1137944CE | M, DREAMWISP NSL, MED CON, W/HGR, CE |
| 1133378RTL | DreamWear Face Mask Kit, M. Cushion |  | 1137944RC | M, DREAMWISP NSL, MED CON, W/HGR, RC |
| 1133430RTL | S, DREAMWEAR FULL CUSHION, RP-RTL |  | 1137945AP | L, DREAMWISP NSL, MED CON, W/HGR, AP |
| 1133431RTL | M, DREAMWEAR FULL CUSHION, RP-RTL |  | 1137945CE | L, DREAMWISP NSL, MED CON, W/HGR, CE |
| 1133432RTL | DreamWear Full Face Cushion, Large |  | 1137945RC | L, DREAMWISP NSL, MED CON, W/HGR, RC |
| 1133433RTL | DreamWear Full Face Cushion, Medium |  | 1137946AP | XL, DREAMWISP NSL, MED CON, W/HGR, AP |
|  |  |  | 1137946CE | XL, DREAMWISP NSL, MED CON, W/HGR, CE |
|  |  |  | 1137946RC | XL, DREAMWISP NSL, MED CON, W/HGR, RC |
| Therapy Mask 3100 NC/SP |  |  | 1137916RTL | FITPK DREAMWISP NSL, MED CON W/HGR -RTL |
| 1144608 | Therapy Msk 3100 NC, fitpack, w/HGR |  |  |  |
| 1144610 | Therapy Msk 3100 SP, fitpack, w/HGR |  |  |  |
| 1145038 | Therapy Msk 3100 NC, fitpack, w/HGR |  |  |  |
| 1145039 | Therapy Mask 3100 NC, fitpack, w/HGR |  |  |  |
| 1145040 | Therapy Mask 3100 NC, fitpack, w/HGR |  |  |  |
| 1145041 | Therapy Mask 3100 NC, fitpack, w/HGR |  |  |  |
| 1145044 | XS, Therapy Mask 3100 NC, w/oHGR, GBL |  |  |  |
| 1145045 | S, Therapy Mask 3100 NC, w/oHGR, GBL |  |  |  |
| 1145046 | M, Therapy Mask 3100 NC, w/oHGR, GBL |  |  |  |
| 1145047 | L, Therapy Mask 3100 NC, w/oHGR, GBL |  |  |  |
| 1145048 | MW, Therapy Mask 3100 NC, w/oHGR |  |  |  |
| 1145049 | Therapy Msk 3100 SP, fitpack, w/HGR, |  |  |  |
| 1145050 | Therapy Mask 3100 SP, fitpack, w/HGR |  |  |  |
| 1145051 | Therapy Mask 3100 SP, fitpack, w/HGR |  |  |  |
| 1145052 | Therapy Mask 3100 SP, fitpack, w/HGR |  |  |  |
| 1145055 | XS, Therapy Mask 3100 SP, w/oHGR, GBL |  |  |  |
| 1145056 | S, Therapy Mask 3100 SP, w/oHGR, GBL |  |  |  |
| 1145057 | M, Therapy Mask 3100 SP, w/oHGR, GBL |  |  |  |
| 1145058 | L, Therapy Mask 3100 SP, w/oHGR, GBL |  |  |  |
| 1145059 | MW, Therapy Mask 3100 SP, w/oHGR |  |  |  |
| 1145060 | Therapy Mask 3100 NC-SP, demopack |  |  |  |
| 1145061 | Therapy Msk 3100 NC, lab fitpack, single |  |  |  |
| 1145062 | Therapy Msk 3100 SP, lab fitpack, single |  |  |  |
| 1145063 | Therapy Mask 3100 NC-SP, lab combo |  |  |  |
| 1145064 | XS, Therapy Mask 3100 NC, cushion, RP |  |  |  |
| 1145065 | S, Therapy Mask 3100 NC, cushion, RP |  |  |  |
| 1145066 | M, Therapy Mask 3100 NC, cushion, RP |  |  |  |
| 1145067 | L, Therapy Mask 3100 NC, cushion, RP |  |  |  |
| 1145068 | MW, Therapy Mask 3100 NC, cushion, RP |  |  |  |
| 1145069 | XS, Therapy Mask 3100 SP, cushion, RP |  |  |  |
| 1145070 | S, Therapy Mask 3100 SP, cushion, RP |  |  |  |
| 1145071 | M, Therapy Mask 3100 SP, cushion, RP |  |  |  |
| 1145072 | L, Therapy Mask 3100 SP, cushion, RP |  |  |  |
| 1145073 | MW, Therapy Mask 3100 SP, cushion, RP |  |  |  |
| 1145074 | Therapy Mask 3100 NC-SP, frame, RP |  |  |  |
| 1145459 | XS, Therapy Mask 3100 NC, w/HGR, GBL |  |  |  |
| 1145460 | S, Therapy Mask 3100 NC, w/HGR, GBL |  |  |  |
| 1145461 | M, Therapy Mask 3100 NC, w/HGR, GBL |  |  |  |
| 1145462 | MW, Therapy Mask 3100 NC, w/HGR, GBL |  |  |  |
| 1145463 | L, Therapy Mask 3100 NC, w/HGR, GBL |  |  |  |
| 1145464 | XS, Therapy Mask 3100 SP, w/HGR, GBL |  |  |  |
| 1145465 | S, Therapy Mask 3100 SP, w/HGR, GBL |  |  |  |
| 1145466 | M, Therapy Mask 3100 SP, w/HGR, GBL |  |  |  |
| 1145467 | MW, Therapy Mask 3100 SP, w/HGR, GBL |  |  |  |
| 1145478 | L, Therapy Mask 3100 SP, w/HGR, GBL |  |  |  |
| 1145916 | Therapy Mask 3100 NC, lab fitpack, GBL |  |  |  |
| 1145917 | Therapy Mask 3100 SP, lab fitpack, GBL |  |  |  |

FIELD SAFETY NOTICE BUSINESS RESPONSE FORM

Reference: 2022-CC-SCR-001

Instructions:   Complete and return all pages of this form within 30 days and return it to Philips Respironics:

safetynoticeuki@philips.com

Customer Name:		_________________________________________________________________________

Street Address: 									_____________________________

City/Postal Code/Country: 					________________________________________________

We acknowledge:

- Receipt and understanding of the accompanying DME/Distributor/Prescriber Urgent Field Safety Notice and confirm that the notice has been provided to all those who need to be aware within your organization and to any organizations where the affected devices have been transferred.

- Receipt of the patient version of the Field Safety Notice and “Facts and Questions” document and confirm that these documents have been provided to all patients using the affected devices.

Name of person completing this form:

Signature: 						_______________________________________________________

Printed Name: 					______________________________________________________________

Title: 							_______________________________________________________

Telephone Number: 				______________________________________________________________

Email Address: 					______________________________________________________________

Date (DD/MM/YYYY): 			____________________________________________________________________